Overview

| Purpose | This document provides guidance and best practices for completing the Fields in the MedWatch 3500A Mandatory Report that is submitted to the FDA. |
| --- | --- |
| Scope | This document applies to Medical Device Reporting (MDR) regulatory reports being submitted to the FDA per 054-WI198 Medical Device Reporting Decision and Regulatory Reporting – United States. - Out of Scope: Outside of the US (OUS) Reporting |
| Roles | Operating Unit (OU) Complaint Unit |

Table of Contents

Process Overview	2

Process Detail	2

1.	Processing Considerations	2

2.	3500A Entries	3

3.	eMDR Submission Status	22

Appendix A – Literature Specific Guidance	22

Appendix B – User Facility and Consumer (Voluntary) MedWatch	24

Appendix C – Event No Longer Qualifies for Reporting	25

Appendix D – Decision Tree for Supplemental Report Type	26

Appendix E – Medtronic Legal Statement	27

Appendix F – Manufacturing Site Information Management in the Global Complaint Handling (GCH) System (Sections D3 and G8)	27

References	29

Process Overview

- Submitted Regulatory Report

- Reportable

- Complaints

Process Detail

# Processing Considerations

The following section includes processing considerations when completing multiple sections of the 3500A form. Information specific to each individual field of the 3500A form is contained in Section 2 of this document.

Review the complaint file to ensure alignment with all fields in the 3500A Form. Note that some fields will auto-populate onto the form; however, if there is information that did not auto-populate, it will need to be manually entered.

| Unknown Information | For any field if it is unknown, leave it blank and follow 027-WI185, Intake, Assessment, and Complaint Determination and 054-WI198, Medical Device Reporting Decision and Regulatory Reporting – United States to complete any necessary follow-up in accordance with 027-WI188, Customer Communication. |
| --- | --- |
| Date Fields | These fields require a complete date (day, month, year). If the date is unknown, follow the guidance below based on available information: - Day is Unknown: Provide known month and year, and use the first day of the month for the day (e.g. 01-MAY-20XX). - Month is Unknown: Provide known day and year, and use the first month of the year for the month (e.g. DD-JAN-20XX) - Year is Unknown: Provide best estimate based on information provided. Note: Enter dates in the date format that is indicated in your complaint handling system. If the system’s date format is different than what the FDA requests, the complaint system automatically puts the date into the correct formation prior to submission to the FDA. |
| Complaints Submitted to the FDA by the Patient, User-Facility, etc. | Refer to Appendix B User Facility and Consumer (Voluntary MedWatch) |
| Changes to FDA Reportability | When an event was reported but additional information reveals it no longer qualifies for FDA reporting, refer to Appendix C Event No Longer Qualifies for Reporting for processing information. |
| Foreign Incident/Event | If submitting a 3500A regarding a foreign incident/event, local privacy restrictions may prevent the inclusion of some patient information. For country specific privacy requirements refer to 027-WI185. |
| Follow Up/Supplemental Reports | - Should not repeat material that was submitted in the initial report but should only provide additional or corrected information from the previously reported event. Note: Exception for drugs and biologics (refer to Section C). - If it is determined that information relevant to the complaint was not sent in a previous FDA report, refer to Appendix D, Decision Tree for Supplemental Report Type for the decision tree to determine the supplemental type, additional information versus correction. |
| Additional Guidance | Information submitted to the FDA becomes publicly available through the FDA’s website. Ensure the RR content is: - Written as clearly as possible, in the past tense, using proper grammar and punctuation, and free of typographical errors. - Does not contain personal identifying or sensitive information such as naming the patient, physician, institution, health care providers, etc - Does not contain inflammatory or defamatory language - Provides accurate and factual information |

# 3500A Entries

The guidance below is based on the 3500A form

| Field | Guidance |
| --- | --- |
| SECTION A | PATIENT INFORMATION |
| A1: Patient Identifier | Provide the patient’s initials or some other identification code number that will allow both the submitter and the initial reporter (if different) to locate the case if contacted for follow-up. For some complaint units, this will auto fill. - Do not use the patient’s name, medical record number or social security number. - If patient initials are unavailable and the patient was enrolled in a clinical study, the patient’s clinical study ID may be added to this field. For no patient involvement, enter “none.” If a country’s privacy laws do not permit the entry of patient information on the regulatory report, this field can be left blank. See 027-WI185 for a list of patient privacy regulations. |
| A2: Age or Date of Birth | Enter the patient’s birth date, if known, or the patient’s age in days, months or years at the time of event onset. - If exact age is unknown, provide the most precise information available. - If age is less than one year, provide months - If age is less than one month, provide days |
| A3: Patient Sex Gender | Make one selection for patient sex. Enter the patient’s sex at birth (the sex that a person was assigned at birth). Male; A person who belongs to the sex that normally produces sperm. The term is used to indicate biological sex distinctions, or cultural gender role distinctions, or both.. Female; A person who belongs to the sex that normally produces ova. The term is used to indicate biological sex distinctions, or cultural gender role distinctions, or both.. The following options are not required on the 3500A form: Undifferentiated Sex; Sex could not be determined; not uniquely defined; undifferentiated.. Prefer Not to Disclose; A response indicating that an individual prefers or preferred not to answer.. Our complaint handling systems may be used to document this information; however it is not required nor part of the 3500(A) Form: Make one selection for the patient gender. Enter the patient’s current gender (how the patient identifies). Male Gender; An individual who reports belonging to the cultural gender role distinction of male.. Female Gender; An individual who reports belonging to the cultural gender role distinction of female. Female-to-Male Transsexual; A person who was assigned to the female gender at birth based on physical characteristics but who self-identifies psychologically and emotionally as male.. Male-to-Female Transsexual; A person who was assigned to the male gender at birth based on physical characteristics but who self-identifies psychologically and emotionally as female.. Other Gender; A response indicating that an individual identifies with a gender category not included in the existing list of values. Response Declined; Used to indicate when a respondent makes a decision to not answer a question.. |
| A4: Weight | Enter patient weight in pounds(lbs.) or kilograms (kgs) at the time of event onset. |
| A5: Ethnicity | Indicate the ethnicity of the patient. Choose only one response. |
| A6: Race | Indicate the race of the patient as reported by the patient. You may choose multiple answers for race. |
| SECTION B | ADVERSE EVENT OR PRODUCT PROBLEM |
| B1: Type of Report | Select the appropriate option: Select adverse event if reporting as a death or serious injury with no associated reportable malfunction. Select product problem for malfunctions that could lead to serious injury or death if they were to recur. Select both if the product problem may have caused or contributed to the adverse event, regardless of the severity of the adverse event. Refer to 054-WI198 for serious injury and malfunction definitions. |
| B2: Outcomes Attributed to Adverse Event | Indicate ALL that apply to the reported event: Death: Check if death was an outcome of the adverse event, or if the cause of the death is unknown. Include the date of death. NOTE: If the date of death is unknown, provide the best estimate and leave a note in the H11. DO NOT check if: - The patient died while using a medical product, but there was no suspected association between the death and the use of the product. Example: A fetus is aborted because of a congenital anomaly or is miscarried. - Life-threatening: Check if suspected that: The patient was at substantial risk of dying at the time of the adverse event. Use, or continued use, of the device might have resulted in the death of the patient. - Hospitalization (initial or prolonged): Check if admission to the hospital or prolongation of hospitalization was a result of the adverse event. DO NOT check if: - A patient in the hospital received a medical product and subsequently developed an otherwise nonserious adverse event, UNLESS the adverse event prolonged the hospital stay. - An emergency room visit did not result in admission to the hospital. However, the event should be evaluated for one of the other serious outcomes (e.g., life-threatening; required intervention to prevent permanent impairment or damage; other serious (medically important event). DO check if: - A patient is admitted to the hospital for one or more days, even if released on the same day. - An emergency room visit results in admission to the hospital. - Disability or Permanent Damage: Select if the adverse event resulted in a substantial disruption of a person’s ability to conduct normal life functions. - Congenital Anomaly/Birth Defect: Select if suspected that exposure to a medical product prior to conception or during pregnancy may have resulted in an adverse outcome in the child. - Required Intervention to Prevent Permanent Impairment/ Damage: If either situation may be due to the use of a medical device and medical or surgical intervention was necessary to: Preclude permanent impairment of a body function. Prevent permanent damage to a body structure. - Other Serious (Important Medical Events): The event is a serious injury and no other boxes from B2 apply to the situation. |
| B3: Date of Event | Provide the actual or best estimate of the date of first onset of the event. Note: For literature events, record the earliest publication date of the article as the date of the event. The earliest publication date may include the date the article was first electronically published, the “e-pub” date. |
| B4: Date of this Report | This is the date when the RR is processed and submitted. This date auto-populates and manual entry is not required. |
| B5: Describe Event or Problem | The Event Description in B5 should match the information as written in the event complaint file. Any event information that is related should be included in the narrative regardless of reportability. This also applies when multiple complaints are captured in the same complaint record and are either not reportable or are reportable on a different product. Reference 027-WI185 Intake, Assessment and Complaint Determination for further details. Only remove items that are administrative e.g., internal dates, source document, and report source tags. |
| B6: Relevant Tests/Laboratory Data, Including Dates | See 027-WI185 for examples. Tests/Data that are associated to the reported event. - **Note: Do not repeat information already provided in the event narrative. |
| B7: Other Relevant History, Including Preexisting Medical Conditions | See 027-WI185 for examples. Information regarding the patient’s medical history, which is related to the event, or preexisting medical conditions that are connected to the patient’s condition. Include any conditions or significant history (including procedures) that is provided in the incoming information. Do not enter internal Medtronic codes or an internal Medtronic review of data logs which results in adding info/diagnoses based on that kind of information. **NOTE: Do not repeat information already provided in the event narrative. |
| SECTION C | DRUG/TISSUE/BIOLOGIC REPORTING |
|  | Section C of the 3500A is filled out when there is a serious adverse experience that is related to the drug portion of the product. G6 will also have 15 days selected. Exception: Follow-up reports on drugs and biologics, including human cells, tissues, and cellular and tissue-based products HCT/Ps, should contain information that was submitted in the original report if the information is still correct. |
| C1: Name, Strength, Manufacturer/Compounder | Space provided to identify two products. Use the trade name as marketed (e.g., if there is a Structured Product Labeling (SPL) for the product, then use the Product name and/or Active ingredient as it appears in SPL; If there is no SPL and a label is attached to the report, then use the same naming convention as in the attached label). If unknown or if no trade name, use the generic name (with the manufacturer or labeler’s name, if known). For foreign reports, use both the foreign trade name and the U.S. generic name. For (HCT/Ps), provide the common name of the HCT/P and the name of the manufacturer. You can also indicate if the HCT/P has a proprietary or trade name, in addition to the common name of the HCT/P (e.g., Brand A bone chips, or Brand B skin tissue). Examples: Achilles’ tendon, Iliac crest bone, Islet Cells, or human skin. For single-entity devices – leave blank Lot # If known, include the lot number(s) with all product problem reports, or any adverse event report that includes a biologic or medication. Example: This may be master data or may need to be found in the OU specific locations such as factory works. NDC # or Unique ID The National Drug Code (NDC #) is a universal product identifier for human drugs. NDC is a three-segment number; zeros and dashes should be included as they appear on the original manufacturer’s product label and/ or packaging. NDC numbers are particularly useful to the FDA in investigating drug product quality problems. Human Cells, Tissues, and Cellular and Tissue-Based Products (HCT/Ps) must have a Unique ID number to track the product. For reports involving HCT/Ps, this unique ID should be provided in this box. |
| C2: List Medical Products and Treatment Given at the Same time of the Event and Date | List and provide therapy dates for any other medical products (drugs, biologics, including HCT/Ps, or medical devices, etc.) that a patient was using at the time of the event. Do not include products used to treat the event. |
| C3: Dose, Frequency, & Route Used | Describe how the product was used by the patient (e.g., 500 mg QID orally or 10 mg every other day IV). For reports involving overdoses, the amount of product used in the overdose should be listed, NOT the prescribed amount. Example: For CV: Use intracardiac or subcutaneous depending on the product. For single-entity devices – leave blank |
| C4: Treatment/Therapy Start and Stop Dates | Provide the date administration was started (or best estimate) and the date stopped (or best estimate). If no dates are known, an estimated duration is acceptable (e.g., 2 years) or, if therapy was less than one day, then duration is appropriate (e.g., 1 dose or 1 hour for an IV). For human cells, tissues, and cellular and tissue-based products HCT/Ps, provide the date of transplant and if applicable, the date of explant. NOTE: Use C4 to record the date that a diagnostic test was performed for reports that involve an in vitro diagnostic product or radiation therapy. This may be the implant or explant date(s). |
| C5: Diagnosis for Use | Provide the indication for which the product was prescribed or used for in the patient (e.g., infection control, suppression of inflammation, etc.). Some indications are mapped. Information may be found in the Product IFU. |
| C6: Product Type | If the product is Over the Counter (OTC), Compounded, Generic or a Biosimilar, please select the best option that best fits this medical product. Leave blank if no fields apply. |
| C7: Expiration Date | Include ONLY with all product problem reports and events involving human cells, tissues, and tissue and cellular-based products (HCT/Ps). For single-entity devices – leave blank. |
| C8: Event Abated after Use Stopped or Dose Reduced | In addition to checking the appropriate box, provide supporting lab tests and dates, if available, in B6. For single-entity devices – leave blank. |
| C9: Event Reappeared after Reintroduction | In addition to checking the appropriate box, provide supporting lab tests and dates, if available, in B6. For single-entity devices – leave blank. |
| SECTION D | SUSPECT MEDICAL DEVICE D1 – D4 Is considered master data and should auto populate, however it is necessary to ensure it is complete. If there is missing or incorrect information, work within the Operating Unit (OU) Complaint Unit to escalate. |
| D1: Brand Name | The trade or proprietary name of the suspect medical device as used in product labeling or in the catalog (e.g., Flo-Easy Catheter, Reliable Heart Pacemaker). |
| D2a: Common Device Name | The generic or common name of the suspect medical device or a generally descriptive name (e.g., urological catheter, heart pacemaker, patient restraint, etc.). Do not use broad generic terms such as “catheter”, “valve”, “screw”, etc. |
| D2b: Procode | Medical devices are classified under 21 CFR Parts 862-892. The assigned FDA Product Classification Code (Procode). Manufacturers and importers should include the Procode for their devices. If the specific device Procode is unknown, enter the most likely or common Procode for the general device involved. |
| D3: Manufacturer Name, City, and State | Enter the full name, city, and state of the Manufacturer* or Specification Developer (design site) of the suspect medical device. See Appendix F Manufacturing Site Information Management in the Global Complaint Handling (GCH) System (Sections D3 and G8) for explanation on Global Complaint Handling (GCH) system management for this section. If D8 below is ‘Yes’, enter the name, city, and state of the re-processor**. *Contract Manufacturers (e.g. Plexus) are not applicable to this definition of 'Manufacturer'. **When a Medtronic product has been re-processed or re-sterilized, the re-processor is then considered to be the manufacturer. Refer to D7. |
| D4: Model # Catalog # Serial # Lot # Expiration date Unique Identifier (UDI) # | If available, provide all identification numbers associated with the suspect medical device exactly as they appear on the device or device labeling. This includes spaces, hyphens, etc. - Model #: The exact model number found on the device label or accompanying packaging. - Catalog #: The exact number as it appears in the manufacturer’s catalog, device labeling, or accompanying packaging. - Serial #: This number can be found on the device label or accompanying packaging; it is assigned by the manufacturer and should be specific to each device. - Lot #: This number can be found on the label or packaging material. - Expiration date: If available; this date can often be found on the device itself or printed on the accompanying packaging. - Identifier (UDI) #: This number can be found on the device, its label, or accompanying packaging and is located below the barcode. FDA defines the UDI: UDI is composed of two parts: Device Identifier (DI) = GTIN Product Identifier(s) (PI) that can include the following: Lot Number Serial number Expiration Date Date of Manufacture Version number EXAMPLE UDI: (01) 12345678901234 (17) 140102 (11) 100102 (10) A1234 (21) 1234 Record all numbers, letters, parentheses, and symbols included in the UDI #. |
| D5: Operator of Device | Indicate the type (not the name) of person operating or using the suspect medical device on the patient at the time of the event as follows: - Health Professional = physician, nurse, respiratory therapist, etc. - Lay Person/Patient/Consumer = person being treated, parent/ spouse/friend of the patient. - Other = nurse’s aide, orderly, etc. In the case of an implant, it is the physician until closure. |
| D6a: If Implanted, Give Date | For medical devices that are implanted in the patient, provide the implant date or best estimate. |
| D6b: If Explanted, Give Date | If an implanted device was removed from the patient, provide the explant date or best estimate. |
| D7a Is this a Single-use Device that was Reprocessed and Reused on a Patient? | Indicate “Yes” or “No.” If the original equipment manufacturer (OEM) is unable to determine if their single use device was reprocessed and reused on a patient, then the OEM should enter ‘UNK’ in H8. In H11 (Additional Manufacturer Narrative), describe the efforts made to obtain the information and any responses. Note: This field is applicable to the re-processor and not Medtronic. |
| D7b: If yes, Enter Name and Address of Re-processor | Enter the name and address of the *re-processor of the single-use device. *When a Medtronic product has been re-processed or re-sterilized, the re-processor is then considered to be the manufacturer. |
| D8: Was this device ever serviced by a Third Party? | Indicate yes, no, or unknown |
| D9: Device Available for Evaluation? | Indicate Y/N whether the device is available for evaluation by the manufacturer. Indicate the date that the device was received at the Medtronic site where analysis can be conducted or the date the product was otherwise made available to Medtronic for analysis. |
| D10: Concomitant Medical Products and Therapy Dates | See 054-WI198 for definition of Concomitant Product. If it is reported that the patient has additional medical products in-use or active at the time of the event that are relevant to the complaint, list the product names and therapy/implant dates for the other medical products (example: drugs, biologics; human cells, tissues, and cellular tissue-based products). This may include other implantable devices that were also impacted by the reported event. Any product that may have caused or contributed to the event should also be reported on a separate 3500A form for that product (unless system reporting applies). Do not include products used to treat the event or ancillary products not associated with the event such as anesthesia equipment, surgical monitors, etc. When the D10 field cannot accommodate all the concomitant products, the remainder of the concomitant products should be placed in H11. Be as specific as possible based on the information available. Follow-up is not required to determine the exact product type. Format suggestions: Model with suffix (e.g., length, if known), device type abbreviations: (e.g., IPG, CRTD, ICM, Lead, Programmer, RF head); therapy date (e.g., implant/explant date). |
| SECTION E | INITIAL REPORTER |
| E1: Initial Reporter Name and Address | Provide the name, mailing address, phone number, and email address of the person who initially reported the adverse event, and who can be contacted to provide information on the event if follow-up is necessary When there is an indication that more than one initial reporter (IR) exists as in the case of a patient alleging a complaint and receipt of a complaint from a Health Care Professional (HCP) select the HCP as the IR. NOTES: - To determine the IR, consider the event in its totality against the dates and known timelines. If information is received that qualifies as a non-reportable complaint, the product is later received and upon analysis the findings support a change to a reportable complaint, the IR would be the person that submitted the information qualifying the complaint as not reportable. If no information was previously received, the product was returned and analysis findings support a reportable complaint, the IR would be Medtronic. - HCP/Physician: Enter the physician's address and contact information in this field. The physician’s facility address may also be used. Proper nouns (ex: facility names) and addresses do not need to be translated into English. - Patient/family/friend: Enter the appropriate information, including the address, of the person who is reporting in this field. For example, if it is a patient’s family member, use the family member information. - Internal (MDT) Initial reporters: If no other Initial Reporter exists, use the Medtronic Rep name along with a Medtronic address and telephone number (not the individual’s personal address or personal phone). - Events occurring in Japan: where the initial reporter is a healthcare professional, patient, or other consumer, the name should not be included due to privacy laws. - Literature events, this field will contain the person listed as the correspondence contact for the article. If a correspondence contact is not listed, use the author. For multiple authors, use the lead author or first author. - For the Diabetes Operating Unit, in the case of an Initial Reporter that resides outside the U.S., see 10702059DOC – MDR Submission in SmartSolve Appendix B- International Contact List. |
| E2: Health Professional | Select Yes if the initial reporter is a health professional (e.g., physician, pharmacist, nurse, etc.) if not, select No. |
| E3: Occupation | Indicate the initial reporter’s occupation. If the complaint handling system’s interface for the 3500A has pre-populated field options, select the most appropriate. If the initial reporter is a healthcare professional, indicate the type of healthcare professional as permitted by the system (e.g., physician's assistant), and include any specialty if known (e.g., cardiology). If the initial reporter is not a healthcare professional, select the most appropriate option permitted in the system. |
| E4: Initial Reporter Also Sent Report to FDA | Indicate whether the initial reporter submitted a notification of this event. See Appendix B for completing a 3500A from a user report. |
| SECTION F | FOR USE BY USER FACILTY OR IMPORTER See Appendix B |
| SECTION G | ALL MANUFACTURERS |
| G1: Contact Office (and manufacturing site for devices) or Compounding Outsourcing Facility | This information is populated by each OU complaint unit. It is typically the information of the OU complaint Unit director who may be contacted if the FDA needs additional information. |
| G2: Report Source | Select the box(es) that most accurately describe(s) the initial (first) source of information leading to Medtronic becoming aware of an event. When submitting a supplemental report, additional boxes may be selected to reflect the source of the new information if that source differs from the initial report source, if applicable. Note: If more than one source informs us independently, more than one selection may be used. - Foreign: if the reported event occurred outside of the U.S. Example: If the event occurred OUS, and there is an HCP listed as the source, select Foreign and HCP. - Study: Post marketing, clinical trial, surveillance, or other study that involves a systematic collection of adverse events from a protocol designed specifically to investigate product safety. - Literature: If the report source is the scientific literature or an unpublished manuscript. Example: If a Health Professional reviews literature that indicates an adverse event and sends it to us for awareness, the original source is (only) literature. - Consumer (including attorneys, patient/family/friend/advocate): If the patient/family/friend/advocate/attorney was the only source, this box should be checked. - Health Professional: Physician, pharmacist, nurse, etc. - User Facility: User facility should be checked if the manufacturer received the report from the MDR contact in a user facility as identified in section F. Example: when a health care facility, submits a User Report (MedWatch) to the FDA who forwards the report to Medtronic. - Company Representative: Check if a company representative reported the event. Example: When analysis findings support there is now a reportable event (where prior to analysis a complaint was not known), the company representative will be selected. - Distributor/Importer: Check if received from the distributor/importer of the suspect product. The health professional or other reporter should be listed as the initial reporter on the front page of the form. - Other: Any source not covered by the previous categories. Other may also be used to identify when the source is another manufacturer. Example: Business Partner for things reported through safety data exchange agreements. Example: MedWatch Report when the information was submitted to the FDA by a patient or patient advocate. |
| G3: Date received by manufacturer | This date is the aware date (typically auto populated). This date is required for all reports (initial and supplemental). For Follow-up reports, use the date that the follow-up information was received. |
| G4: For use by all Manufacturers of Drug, Device, Biological products [including Cell, Tissue, Cellular and Tissue-Based products (HCT/P)] and Combination products. | This section is mostly Master data. - Confirm the PMA/510(k) number. If there are multiple PMA or 510(k) numbers available and it is difficult to tell which is the most applicable, use the most recent. Where needed, resources for verifying PMA/510(k) information include the FDA GUDID database, the FDA Medical Device Establishment Registration Device Listing database, as well as local regulatory affairs teams. If a 510(k)/PMA is not available (example: not approved within the US, unknown general product or product has an exempt Procode), then the field may be left blank. - Additionally, if appropriate, use this field to enter the abbreviated new drug application number, new drug application number, investigational new drug application number or biologics application number. - Check box for Combination Product if applicable. |
| G5: Give IND/Pre ANDA Protocol # | Not applicable for most teams, only used for entering the applicable protocol number for investigation of new drug safety reports. |
| G6: Type of Report | Check all boxes that apply to reported event. See 054-WI198 for definitions: - 5-day - 7-day - 15-day - 30-day - Periodic - Initial: Check if the report is the first submission of a manufacturer report. For devices, this is the 30-day report. - Follow-up: Check if the report is a follow-up to a previously submitted report. |
| G7 Adverse Event Term(s) | For use by drug and biologic, including human cell, tissue, and cellular and tissue-based product (HCT/P), [manufacturers only] Include a list of adverse event terms that most accurately characterize the adverse event described in narrative format in B5. Terms should be listed with the most relevant term(s) first. The terminology may be an accepted standard (e.g., MEDDRA or WHOART), a verbatim term, or the manufacturer’s own terminology. |
| G8: Manufacturer Report Number | This number uniquely identifies each individual report submitted by a manufacturer or importer. This number consists of the following three parts: (1) The FDA registration number for the manufacturing site of the reported device, or the registration number for the importer. If the manufacturing site or the importer does not have an establishment registration number, we will assign a temporary MDR reporting number until the site is registered in accordance with part 807 of this chapter. We will inform the manufacturer or importer of the temporary MDR reporting number; (2) The four-digit calendar year in which the report is submitted; and (3) The five-digit sequence number of the reports submitted during the year, starting with 00001. (For example, the complete number will appear as follows: 1234567-2011-00001.) For Manufacturers and Importers: The manufacturing report number will auto-populate. See Appendix F - Manufacturing Site Information Management in the Global Complaint Handling (GCH) System (Sections D3 and G8) for details of part 1 of the manufacturing report number in the Global Complaint Handling (GCH) system. For drug and biologic manufacturers: The manufacturer report number is the number the manufacturer chooses to uniquely identify the report and should conform to any applicable regulations or guidance’s. If submitting a follow-up to a report originally obtained from FDA through a MedWatch to Manufacturer program transmission of a serious direct report, check the “Other” box in G3 and enter the FDA-assigned report number there. For human cell, tissue, and cellular and tissue-based product (HCT/P) manufacturers: The report number should consist of three numbers separated by dashes. The first number will be the 10-digit FDA Establishment Identifier (FEI) number, which was assigned as part of the Human Cells and Tissue Establishment Registration (HCTERS). The second number should be the year that the report is being submitted. The last number should be a consecutive 5-digit number for each report filed during the year by the manufacturer. Example: 1234567890- 2016-00005. |
| SECTION H | DEVICE MANUFACTURER ONLY |
| H1: Type of Reportable Event | Select the box with the highest level of severity that applies to this report. Refer to 054-WI198 for definitions: - Death - Serious injury - Malfunction - Summary Report Examples: - If the reason for reporting is death, check “Death”, regardless, if there was also a serious injury or a malfunction associated with the event. - If the reason for reporting is serious injury but no death, check “Serious Injury”, regardless, if there was also a malfunction associated with the event. - If the reason for reporting is MDR-reportable malfunction and there is no death or serious injury associated with this report, check “Malfunction”. |
| H2: If Follow-up, What Type? | Check all box(es) that apply: - Correction - Additional information - Device evaluation See Appendix C, Decision Tree for Supplemental Report Type for additional guidance on correction versus additional information reports. |
| H3: Device Evaluated by Manufacturer? | Check the dropdown marked “Yes” if the suspect or related device was evaluated. Provide a summary of the evaluation in H11: Additional Manufacturer Narrative. If an evaluation of a returned suspect or related medical device was NOT conducted, check the dropdown marked No. Use appropriate device evaluation code in H6 to indicate device was not evaluated, or, provide a justification in H11 explaining why it was not. |
| H4: Device Manufacture Date | Enter the manufacture date of the medical device. If the full date is unknown, use the best estimate and note this in H11. This information may be part of Master Data for some OUs. If it is missing or unclear, see OU specific guidelines for direction on gathering this information from internal sources (e.g., Factory Works, MTCS, DHR, etc.). |
| H5: Labeled for Single Use | Indicate whether the device was labeled for single use. Ensure H5 does not conflict with D7 and H8. This information may be part of Master Data and auto fill based on the product. See OU specific guidelines for directions if missing. |
| H6: Adverse Event Problem | Enter all code types. Some Complaint Handling Systems may have entry of these codes in various places. There must be at least one of each code type provided. Patient Health Effect Clinical Code(s): E code Device Problem Code(s): A code Type of Investigation Code(s): B code Investigation Finding Code(s): C code Investigation Conclusion Code(s): D code Health Effect Impact Code(s): F code Component Code(s): G code Note: Codes should not be removed from H6 even if not reportable or duplicate. If a code has been submitted on a prior report for this product, then the repeated code can be removed. |
| H7: If Remedial Action Initiated, Check Type | Do not check any section within H7 unless information has been received that indicates the involved product lot/serial are involved in a recall or other field corrective action (FCA). - Check “recall” and/or “notification” depending on the event circumstances. Ensure the correction and removal number is entered in H9. Most of these terms are defined or further explained in the Act 21 USC 360h or in the FDA regulations 21 CFR Parts 7, 803 and 806 concerning remedial action. - If Medtronic determines to initiate remedial action for a device problem after an initial MDR report has been submitted, a supplemental MDR report must be submitted capturing the remedial action taken in response to the initially reported problem. NOTE: For events that were submitted on the 806 report, a supplemental report is not required. - The supplemental MDR should be filed within 30 days of the decision to initiate remedial action. If remedial action is not initiated in response to a reported event, H7 and H9 are to be left blank. - When additional events are identified as associated to 5-day reports that alerted Medtronic to initiate remedial action, the additional events will be submitted as 30-day reports. The complaints that initiated a recall are often listed in associated Issue Impact Assessments. |
| H8: Usage of Device | Indicate whether the use of the medical device was the initial use, reuse, or unknown. If a manufacturer receives an adverse event report that indicates that the event was caused or contributed to by reuse of a single use device they manufactured, this field is to be marked and the facts of the firm’s investigation provided with an explanation of how the reuse of the product contributed to the outcome. The appropriate manufacturer Result codes for reuse are also to be entered into H6. |
| H9: If action reported to FDA under 21 USC 360i(g), list correction/removal reporting number | If the complaint is associated with a recall (as indicated in section H7), enter the correction/removal reporting number or equivalent here. NOTE: - The 806 correction/removal reporting number consists of the registration number for the responsible firm, the date of the report using MM/DD/YY format, a 3-digit sequential number for each report made on that date (i.e., 001, 002, 003, etc.), and the report type designation (e.g., “C” for a report of correction or “R” for a report of removal). - The FDA recall number consists of Z-XXXX-XXXX. - FCA that have not been reported on an 806 form, do not require input in this field |
| H10: Related Report Number | When there are multiple products within the same complaint record that are reportable to the FDA, this field is to include the Manufacturer Regulatory Report number (G8) for the other reportable products that are related to the event. This may also include a reference to a duplicate RR if multiple RRs were submitted in error. NOTES: - MedWatch User Facility (UF) or assigned MedWatch number. - Once one FDA Regulatory Report (RR) has been submitted within the event, the complaint system may automatically add RR numbers to FDA RRs generated for other products in the same event. Example: Product #1 RR is submitted. Product #2 will contain the RR # for product #1. Product #3 will contain the RR #’s for products, 1 and 2. - A Supplemental Report solely to include a related RR number is not necessary. |
| H11: Additional Manufacturer Narrative | Enter any additional information, evaluation, or clarification of data presented in previous sections. Do not duplicate information that has already been provided elsewhere. This section may include: - Concomitant products overflow-suggested format: Continuation of D10: Brand Name ProductID Lot/Serial Number (Implant: XYZ)(Explant: XYZ) - Correction or Clarification Statements: List fields that where the supplemental report information has changed or was corrected. The corrected information itself should be reported in the specific sections of the 3500A. - Patient Privacy statement (suggested verbiage): “Select patient information cannot be provided due to regional privacy regulations." - If there are duplicate or non-reportable codes in H6 (suggested verbiage): “H6: The codes present in H6 correspond to components/products that comprise the reported event.” - Analysis Summary (where applicable) - Investigation Summary (where applicable) - Similar or Same product notification for those OUs that have this in their process. Suggested format(s): - Similar Product “Please note that the involved device is not marketed in the United States; however, it is similar to the United States marketed device (Brand Name ProductID). - Same Product: “Please note that the involved device is not marketed in the United States; however, the device is deemed the same as the United States marketed device (Brand Name ProductID).” - Suggested formatting for reports sent that involve system reporting: - This is a system report. The section D information is for the primary device, which was in use with the following: Brand Name [Brand Name] Product ID [Product ID] (Lot: [Lot#] Serial: [Serial #]). - Section F input for User report receipts. See Appendix A – User Facility and Consumer (Voluntary) Medwatch - Medtronic Legal statement (auto-populates after submission. See Appendix D – Medtronic Legal Statement for full copy) - Variance 5 (Diabetes) “This MDR related to the Puerto Rico manufacturing site has been assigned a MedWatch number from the Medtronic MiniMed Northridge site, per Variance 5.” Literature Events: Foreign language articles should have full citations provided. Example of literature citation: AUTHORS. “TITLE”. JOURNAL NAME, ISSUE#, YEAR PUBLISHED. LITERATURE REFERENCE *Literature reference = Website.PMID:___. Pages:____. Cosgun MS, Cosgun C. “Predictors of shoulder limitations and disability in patients with cardiac implantable electronic devices: Importance of device size.” Journal of Sports Economics, no. 3, 2000, https://doi.org/10.1177/152700250000100304. PMID:12345. Pages:304-307. For Diabetes: Always mark “Additional Mfr Narrative” for initial MDRs and for Supplemental Reports it could be either “Additional Mfr Narrative” or "Corrected Data" or both. |

# eMDR Submission Status

The following indicates the submission status and receipt of the regulatory report by the FDA.

| Status of Completed eMDRs | Definition | Source |
| --- | --- | --- |
| Blank | No eMDR status has been set for this report | - Medtronic |
| Report Submitted | Indicates the report has been sent through the gateway to the FDA. This status will be set by the process linking to the eMDR tracking database. |  |
| ACK1 Receipt | Also referred to as Receipt. Sent by the Electronic Submissions Gateway (ESG) and confirms the submission was successfully received by the ESG. |  |
| ACK2 Received | Indicates an acknowledgement from the FDA that the report has been received and processed by the Center X Radiologic Health (CDRH). This status is set by the batch process linking to the eMDR tracking database |  |
| ACK3 Received Successful | Indicates acknowledgement of receipt by the FDA and the report has been processed by CDRH. This status is set by the batch process linking to the eMDR tracking database Note: If ACK3 is not indicated submit a ticket to Medtronic IT Systems Support. |  |
| eMDR Transmission Error | Indicates errors were encountered, and the report was not successfully processed by the FDA. This status is set by linking to the eMDR tracking database. |  |

Appendix A – Literature Specific Guidance

| Considerations | Action |
| --- | --- |
| For Scientific Literature or Unpublished Manuscript | Attach a copy of the literature article to the Regulatory Report or provide the web address location for the article (refer to the example in H11 for the literature citation format) |
| Copyright Permission | To verify copyright permissions via the Copyright Clearance Center, go to http://rightfind.copyright.com and search by title, or International Standard Book Number (ISBN)/International Standard Serial Number (ISSN). |
| For Drugs and Biologics, including Human Cells or Tissues (HCT) | Complete a separate 3500A form for each identifiable patient in the article or manuscript. |
| Number of Regulatory Reports | For all reportable events identified through literature sources (published or unpublished), when there is limited information about the product and or the patient and follow-up is not possible: - Submit an individual report for each event type if more than one is identified. For example, if the literature source includes information about patient deaths, serious injuries, and a reportable malfunction, submit three individual reports (one for each event type). Submit an individual report for each device identified if the generic names of the products are different. For example, if a literature source identifies the suspect products as catheters and pacemakers, then submit two reports: one for each product. Consider the different event types as noted above. |

Appendix B – User Facility and Consumer (Voluntary) MedWatch

When product malfunctions and/or adverse events are reported directly to the FDA through a:

- 3500A Mandatory MedWatch (by user-facilities, importers, distributors, and manufacturers)

- 3500 Voluntary MedWatch (patient, patient advocate)

The FDA forwards these reports to Medtronic.  The Complaint Handling Unit will submit a 3500A with the specific information reported to the FDA along with the information from the table below. If the complaint unit determines that event does not meet MDR reporting criteria, then no 3500A submission is necessary. If an MDR has already been submitted for the event and there is no additional information requiring a supplemental 3500A, then no additional submission is needed. If the only new information received are updates to fields E4 and G2, a supplemental is not required.

When a User Facility (UF) submits a report to the FDA a UF report number is assigned. Reference this number in H10.

When submitting a 3500A, ensure the correct information is captured in E4 and G2.

| E4 Report sent to the FDA | Select Yes if the User Facility reported the event to the FDA |
| --- | --- |
| G2 Report Source | See examples above in G2 |

Appendix C – Event No Longer Qualifies for Reporting

A regulatory report that has been submitted and later found to not qualify for reporting cannot be retracted.

- Events may be updated as ‘non-reportable’, although the information will remain with the US FDA and may still exist in the public domain. This may be due to the event being incorrectly reported initially or due to new information received indicating the event no longer meets reportability criteria.

- Consult your CQXM management to determine if an event needs to be un-reported. CQXM management may defer this responsibility to analysts. If the decision is to un-report, follow these steps:

Update the complaint handling system as necessary.

Update the follow-up MDR report:

If reported in error, update the following:

B5: Event description/narrative based on the updated information, or statement indicating why this is being updated to non-reportable.

H2 Check Correction

H11 correction B5; Suggested verbiage: “Additional review of the event and/or additional information received shows that there is no information to reasonably suggest that the device in this report may have caused or contributed to a death or serious injury or that the device in this report has malfunctioned.  Therefore, this event did not and does not meet the reporting requirements in 21 CFR 803.”

If new information indicates that the event is not reportable, update the following:

B5: event description/narrative based on the updated information, or statement indicating why this is being updated to non-reportable.

H2 Check Additional Information

H11 add the following text: “Additional information received shows that there is no information to reasonably suggest that the device in this report may have caused or contributed to a death or serious injury or that the device in this report has malfunctioned.  Therefore, this event does not meet the reporting requirements in 21 CFR 803.”

Example:

If a duplicate report was inadvertently submitted, submit a follow-up/supplemental to indicate the event was previously reported under a different regulatory number. No other data (i.e., analysis results would be included.

Appendix D – Decision Tree for Supplemental Report Type

The decision tree below provides guidance when a “correction” versus “additional information” supplemental report is required.

An additional information supplemental regulatory report is submitted when new information that is required on the 3500A report becomes available and it was not known by any Medtronic employee when a previous 3500A report was submitted for that event.

A correction supplemental regulatory report is submitted when information that is required on the 3500A report was available to any Medtronic employee when a previous 3500A report was submitted for that event. This can include information that was missed, incorrectly assessed, or not communicated to the OU Complaint Unit at the time of previous 3500A submission(s).

Appendix E – Medtronic Legal Statement

This legal statement may auto-fill for some complaint handling systems. This statement is not visible to the user while processing the RR. The complaint handling system auto-populates this into H11 when the system submits the report.

“Medtronic submits this report to comply with FDA regulations 21 CFR Parts 4 and 803. Medtronic has made reasonable efforts to provide as much relevant information as is available to the company as of the submission date of this report. This report does not constitute an admission or a conclusion by FDA, Medtronic, or its employees that the device, Medtronic, or its employee caused or contributed to the event described in the report. Any required fields that are unpopulated are blank because the information is currently unknown or unavailable. Medtronic will submit a supplemental report if additional relevant information becomes known.”

Appendix F – Manufacturing Site Information Management in the Global Complaint Handling (GCH) System (Sections D3 and G8)

When completing the 3500A MedWatch, via the Global Complaint Handling (GCH) system, the manufacturing site information, which is used to auto-populate Sections D3 and Part 1 of section G8, is managed within the data of the Reg Report Details tab on the GCH Regulatory Report submission page.

- The Mfg Site ID (also known as the PCode) and FDA Site ID will auto-populate in most instances.  However, there are several scenarios where they may not auto-populate.  The following describes those scenarios, and the steps to ensure the information is complete.

- Manufacturing Site and FDA Site ID Alignment:

| - Scenario | - Reason | - Guidance |
| --- | --- | --- |
| - Mfg Site ID does not populate | - There is not enough available information to populate data from Master Data (i. e. the Product Line Item (PLI) does not populate a Mfg Site ID (PCode)). - Usually occurs when: - Unknown Products - Lot/Serial Number are not available or not valid | - Select the most appropriate FDA Site ID (e.g. the most recently registered Manufacturer or Specification Developer, as mentioned in Section D3) for the product. * |
|  | - Product information does not exist in the Reg Report Details data to auto-populate this section. | - Work with the local SME and/or the OU to determine the appropriate ID(s) for use. * - Submit a GCH IT Ticket that includes the following details (see the example below the table): Brief Description: “GCH – Update Mfg Site Information in the ZCRM_RRNO table” Details: Request to update the ZCRM_RRNO table with the following: Mfg Site ID (PCode from Asset/Master Data) FDA Site ID* Manufacturer Name* Manufacturer City* Manufacturer State* Manufacturer Country* |
| - Mfg Site ID does populate, and FDA Site ID provides more than one option | - Mostly occurs when a Contract Manufacturer manufactures products for more than one OU. This will present multiple options for the 'FDA Site ID' field in the 'Reg Report Details' (i.e. one FDA Site ID for each OU). | - Select the most appropriate FDA Site ID for the product. It is recommended to use the Specification Developer site. * |
| - Mfg Site ID populates, but FDA Site ID is not aligned | - Occurs when product information (specifically, the FDA Site ID) does not exist in the Reg Report Details data to auto-populate this section (e.g. shows with incorrect address or site). | - Work with the local SME and/or the OU to determine the appropriate ID(s) to use. * - Submit a GCH IT Ticket that includes the following details (see the example below the table): - Brief Description: “GCH – Update Mfg Site Information in the ZCRM_RRNO table” - Details: Request to update the ZCRM_RRNO table with the following: - Mfg Site ID (PCode from Asset/Master Data) - FDA Site ID* - Manufacturer Name* - Manufacturer City* - Manufacturer State* - Manufacturer Country* |

Example of GCH IT Ticket Request details

* If needed, work with the local Operating Unit to confirm the correct registered site information for specific products.  Generally, the most appropriate site can be identified by searching in FDA’S Unified Registration Listing System (FURLS).  This can be accessed via the Establishment Registration & Device Listing site (https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRL/rl.cfm).

References

| Number | Title |
| --- | --- |
| 027-WI185 | Intake Assessment and Complaint Determination |
| 027-WI188 | Customer Communication |
| 054-WI198 | Medical Device Reporting Decision and Regulatory Reporting – United States |
| 10702059DOC | MDR Submission in SmartSolve Appendix B- International Contact List |